## Q1
Question 1 - Regarding timing in endoscope reprocessing: | Responses
How many endoscopes are reprocessed at your trust per month? | 1300
Is the time interval between procedure and pre-cleaning recorded? If so, how is this recorded? | no
What is the average time taken to reach a pre-cleaned state? | N/A
Is the time interval between pre-cleaning and manual cleaning recorded? If so, how is this recorded? What is the average time taken to reach manual cleaning? | It is not routinely recorded/monitored, but audited manually. It is taken from the post-procedure clean documentation in patient records and 'decon start' time in the reprocessing centre. Average time in last two audits was 1h48m and 1h49m
Are there timestamps for individual stages of reprocessing? How are timestamps recorded during reprocessing? | yes captured on electronic tracking system in reprocessing centre
Is the duration of manual cleaning recorded? If so, how is this recorded? What is the average duration of manual cleaning? | Through SynergyTrak online Track Trackability portal. 5 minuites each scope
Is endoscope turnaround time monitored? If so, how is this monitored? | This is monitored through ST and Endoscopy clinical team write on the scopes the finished procedure time.
What is the average turnaround time now? What was it 6 months ago? | Now - 1:12 - 6 months ago - 5:31
How are critical time barriers monitored? | Through Trust decon Lead audit. (Manual)
Is there a staff member responsible for time management? If so, who undertakes this responsibility? | No, as this is monitored through SynergyTrak and scopes are turned around ASAP
Has the Trust identified any challenges associated with monitoring elapsed-time? | We currently use a manual process with clinical teams, and electronic system in reprocessing centre. There are challenges in visibility between the two systems, hence the audit.

## Q2
Question 2: Regarding staffing and traceability | Responses
How many staff members are designated for endoscopy reprocessing? | 3 FTE
Are staffing levels for manual cleaning monitored? How many staff members are involved in this process? | Monitored by how many scopes are available to process, through a performance indicator
Are staffing levels for pre-cleaning monitored? How many staff members are involved in the process? | not separately - part of clinic/procedure task and staff workload
How many staff members are involved in the transport process (from pre-clean to manual clean)? | 1 x FTE
Does your trust use an electronic tracking/traceability system? If so, what is the name/type of the system? | Electronic tracking in reprocessing centre. SynergyTrak
How do staff know which scopes should be reprocessed first? | Through the time procedure ends on the QA label and is monitored by SynergyTrak
Would staff be interested in time-lapse indicator labels? | If scanning was set up in each area, this would not be needed, as it would be monitored end to end through SynergyTrak
Has the Trust experienced staffing shortages within endoscope decontamination services during the last 12 months? | No
Does the Trust consider staffing pressures a risk to maintaining reprocessing turnaround times? | Yes

## Q3
Question 3 - Regarding delays in endoscope reprocessing | Responses
Is delay information formally recorded? If so, how is this recorded? | audits reprocessing centre data can be reported
What are the most common causes of delays in reprocessing? | SSL staff have to transport the scopes around the hospital for delivary
During the last 12 months, how many endoscopes required repeat or additional reprocessing due to delays between bedside pre-cleaning and manual cleaning? | N/A
How many scopes, on average, are not reprocessed within one hour per month? What was this figure a year ago? | 526 =39,7% ---- 1 year ago 459 = 33.90%
How many endoscopes are reprocessed out of sequence per month? What was this figure a year ago? | N/A
What is the distance travelled by a scope in-between pre-cleaning and manual cleaning? | N/A
During the last 12 months, how many procedures were delayed due to endoscope unavailability? | This has never as if this had occurred in the last 12 months, this would have been escalated to the operational manager and Deputy Divisional Manager & Access Manager Surgery
During the last 12 months, how many procedures were cancelled due to endoscope unavailability? | This has never occurred as it would have been escalated to the operational manager and Deputy Divisional Manager & Access Manager Surgery if we were cancelling procedures due to scopes not being available.
Does the trust record causes of endoscope unavailability? | Repair trends are monitored including reason for repair
During the last 5 years, has the Trust investigated any incidents where delayed reprocessing contributed to an endoscope reprocessing concern? | No
Does the Trust recognise delayed reprocessing as a patient safety risk within its endoscope management procedures? | yes and is monitored/audited.
Has the Trust estimated the cost associated with endoscope reprocessing? If yes, please provide any documented cost estimates. | Approx £47 per scope

## Q4
Question 4 - Regarding endoscope cleaning practices | Responses
During the last 12 months, how many endoscopes were removed from service due to blocked channels? | X 6 scopes
Does the Trust record incidents involving blocked endoscope channels? If yes, please provide the number of incidents recorded in the last 12 months. | 1 reported via Datix
Which endoscope channel(s) are most commonly reported as blocked? 1. Biopsy channel 2. Suction channel 3. Air/water channel 4. Auxiliary channel 5. Other | Records are not kept with that level of detail
How many endoscopes required repeat manual cleaning during the last 12 months? | None
How many endoscopes failed cleanliness verification tests during the last 12 months? Does the trust record reasons for failed cleanliness tests? | None
Has the trust assessed whether delays between pre-cleaning and manual cleaning contribute to cleaning difficulties or repeat processing? | It is something we explore on a case-by case basis. If there were themes we would explore in more detail.
During the last 3 years, has the Trust investigated any incidents involving: residual soil, channel blockage, inadequate cleaning or endoscope contamination? If yes, how many incidents were recorded? | 5 reported on Datix

## Q5
Question 5 - Regarding documentation | Responses
During the last 3 years, has the Trust received any audit findings relating to endoscope reprocessing compliance? Were any findings related to documentation, traceability, timing or workflow controls? | External audits (e.g. JAG) - nil Local audits - some gaps identified and addressed
Has the Trust implemented corrective actions relating to endoscope reprocessing workflow? | yes
During the last 3 years, how many incident reports related to endoscope reprocessing have been recorded? 1. How many are related to: 1. Delayed reprocessing? 2. Documentation issues? 3. Scope traceability? 1. Cleaning failures? 2. Scope availability? | delayed reprocessing - 7 (early theme within 3yr timeframe, numbers reduced with corrective actions) documentation - 8 (early theme within 3yr timeframe, numbers reduced with corrective actions) scope traceability - 1 cleaning failures - 5 (missed post procedure flush) scope availability - 3 (related to equipment repairs rather than reprocessing )